## Slide 1: Hitting the bull’s eye in proton therapy Imaging patient and beam: overview and recent work at KVI-CART
Peter Dendooven
91st meeting
16 March 2018 - Amsterdam

## Slide 2: Particle beam radiotherapy worldwide and in NL
175 000 patients treated (end of 2016)
20 000 patients treated in 2016
79 facilities in operation
46 facilities under construction
22 facilities under planning
source: www.ptcog.ch
Delft
Groningen
Maastricht
- - - - - - - - - - Groningen
- - - - - Delft, Maastricht
(Amsterdam is undecided)

## Slide 3: UMC Groningen Proton Therapy Center
IBA ProteusPLUS
2 treatment rooms
capacity ~ 600 patients per year
www.umcgroningenptc.nl/en/
start of operations: 22 January 2018
first proton therapy treatment in the Netherlands

## Slide 4: Proton therapy workflow (simplified)
radiation dose prescription
treatment plan
treatment
How do protons behave inside the patient ?
Was the dose delivered according to plan ?
www.raysearchlabs.com
iba-worldwide.com
CT
Notes: Focus on 2 aspects:
How do protons behave inside the patient ?
Needed for good treatment planning.
Was the dose delivered according to plan ?

## Slide 5: Knowing the patient
anatomy
proton stopping power, proton range
Ye Zhang et al., Phys. Med. Biol. 57(2012)1779

## Slide 6: Knowing the patient anatomy
standard: CT

## Slide 7: Knowing the patient anatomy
standard: CT
MRI: under development
D.A. Low, in “Advances in Radiation Oncology”, Eds. Wong, Schultheiss, Radany (2017)
Notes: Why MR ? Under development at least for photon radiotherapy.
better identification of different soft tissues
ability to acquire multiple image datasets that highlight the tissue boundaries, and differentiate different tissue types, even within a specific tumor
Drawback: no direct relation to electron density and average atomic number, the material properties important for proton stopping power
Fig. from D.A. Low, in “Advances in Radiation Oncology”, Eds. Wong, Schultheiss, Radany (2017)

## Slide 8: From anatomy to proton stopping power
CT
3-4% uncertainty in proton range
imposes a safety margin around tumour
better
information at different X-ray energies
information from protons
Notes: main problem: protons interact fundamentally different with matter than photons
Better than CT: measure attenuation at different X-ray energies
high energy, Compton scattering  e
low energy: photo-electric effect + Compton
techniques: DECT and spectral CT

## Slide 9: Different X-ray energies: DECT
dual-energy CT
www.healthcare.siemens.com
Notes: main problem: protons interact fundamentally different with matter than photons
Better than CT: measure attenuation at different X-ray energies
high energy, Compton scattering  e
low energy: photo-electric effect + Compton
techniques: DECT and spectral CT

## Slide 10: Experimental investigation
(DE)CT
at UMCG
proton stopping power
at KVI-CART
high-accuracy (≤ 0.05 mm) range measurements

## Slide 11: Different X-ray energies: DECT
dual-energy CT
100 kV
140 kV
Notes: main problem: protons interact fundamentally different with matter than photons
Better than CT: measure attenuation at different X-ray energies
high energy, Compton scattering  e
low energy: photo-electric effect + Compton
techniques: DECT and spectral CT

## Slide 12: Single energy vs. dual energy CT
tissue-equivalent materials
animal tissues
accuracy of relative stopping power
J.K. van Abbema et al.,Phys. Med. Biol. 60(2015)3825
J.K. van Abbema, PhD Thesis, University of Groningen (2017)
more accurate patient specific tissue proton stopping powers

## Slide 13: Different X-ray energies: spectral CT
Shikhaliev and Fritz PMB 56 (2011) 1905
Notes: How much extra does this bring ? Needs to be determined.

## Slide 14: Stopping power imaging: proton imaging
proton beam
for treatment
head phantom scan; 219 MeV protons
proton beam for
proton radiography
T. Plautz et al, IEEE Trans Med Imag (2014) 33, 875-881
source: U. Schneider, PSI

## Slide 15: Proton radiography at KVI-CART
A.K. Biegun et al. Physica medica 41(2017)141 A. Biegun et al., J. Instrum.11(2016)[C12015
A.K. Biegun et al., IEEE Trans. Med. Imaging 35(2016)1099 A. Biegun et al., Acta Phys. Pol. B 47(2016)329
collaboration with
detector
detector
phantom
Notes: Practical implementation of proton imaging is some time off. Several prototype systems around the world, non in routine clinical use.
Investigate e.g. the effect of angular selection on image quality and requirements needed for the absorption detector.

## Slide 16: Proton therapy workflow (simplified)
radiation dose prescription
treatment plan
treatment
Was the dose delivered according to plan ?
www.raysearchlabs.com
iba-worldwide.com
Notes: Focus on 2 aspects:
How do protons behave inside the patient ?
Needed for good treatment planning.
Was the dose delivered according to plan ?

## Slide 17: In-vivo verification
differences in
anatomy / positioning: “see what you treat”
proton range
if actual treatment situation ≠ treatment planning situation
→ serious dose deposition errors may occur

## Slide 18: Deviation from the treatment plan: an example
before treatment
source: S. Mori en G.T.Y Chen, MGH
after 5 weeks

## Slide 19: Deviation from the treatment plan: an example
before treatment
after 5 weeks
source: S. Mori en G.T.Y Chen, MGH

## Slide 20: See what you treat: X-rays
in-room imaging for daily anatomy / positioning
  X-ray radiography – cone beam CT
Notes: realtime anatomical feedback can be used in several ways:
motion tracking to control irradiation in real time (gating or steering)
calculate the actual dose distribution delivered
This technique is being introduced in photon beam therapy, but is even more appealing for proton therapy, because of the higher sensitivity to differences between the actual situation and the treatment situation.

## Slide 21: See what you treat: MR imaging
real-time anatomical feedback:
motion tracking for real-time irradiation control, adaptive workflow
calculation of the actual dose delivered
S. Crijns and B. Raaymakers, Phys. Med. Biol. 59(2014)3241
B.W. Raaymakers
Notes: realtime anatomical feedback can be used in several ways:
motion tracking to control irradiation in real time (gating or steering)
calculate the actual dose distribution delivered
This technique is being introduced in photon beam therapy, but is even more appealing for proton therapy, because of the higher sensitivity to differences between the actual situation and the treatment situation.

## Slide 22: Real-time MRI-guided proton therapy
B.M. Oborn et al., Med. Phys. 44(2017)e77
possible layout

## Slide 23: Collaboration on MRI-guided proton therapy
degrader + energy
selection system
scanning magnets

## Slide 24: In-vivo dose delivery verification
How ?
measuring secondary radiation:
  positron emission tomography (PET)
  prompt gamma rays
  sound waves
Notes: even if anatomy and positioning are fine, proton range can be wrong
sound waves is less developed than others, PET is used in the clinic in a few centers around the world

## Slide 25: In-vivo dose delivery verification using PET
treatment planning CT
oxygen-15 production
treatment planning CT
potassium-38 production
treatment planning CT
simulated dose
P. Dendooven et al.
principle: production of positron emitters
correlates with radiation dose
Notes: Movies:
IEEE 2012 > Hadron Therapy workshop talk >
Rendering_CT_dose.wmv
Rendering_CT_oxy15.wmv
Status Simulations August 2012 > RadiationDistributions >
Rendering_Pot38Map_CT2705_BU+Vinnie Full CT raw data format from AMIDE_1.wmv

## Slide 26: In situ / beam-on PET - an attractive option
in-situ / beam-on
figure from Zhu X and El Fakhri G, Theranostics 3 (2013) 731
T. Nishio et al., Int. J. Rad. Oncol. Biol. Phys. 76 (2010) 277
National Cancer Center, Kashiwa, Japan
dual-head
PET scanner
proton beam

## Slide 27: In situ / beam-on PET - an attractive option
benefits:
  maximum number of counts
  minimal biological washout
  faster feedback
very short-lived nuclides are relevant
in-situ / beam-on
figure from Zhu X and El Fakhri G, Theranostics 3 (2013) 731

## Slide 28: Nitrogen-12 imaging (half-life = 11 ms)
Buitenhuis et al., Phys. Med. Biol. 62(2017) 4654–4672
Dendooven et al., Phys. Med. Biol. 60(2015) 8923–8947
all
long-lived
short-lived
90 MeV protons
Philips PET detectors
experiments at KVI-CART
PET image profiles
proton beam
Notes: observed shift: 6±3 mm
simulations of realistic scanner and number of protonts: ~1 mm accuracy
same on bone

## Slide 29: Nitrogen-12 imaging (half-life = 11 ms)
90 MeV protons
Philips PET detectors
experiments at KVI-CART
Buitenhuis et al., Phys. Med. Biol. 62(2017) 4654–4672
Dendooven et al., Phys. Med. Biol. 60(2015) 8923–8947
all
long-lived
short-lived
PET image profiles
proton beam
millimeter range accuracy on 50 ms time scale
Notes: observed shift: 6±3 mm
simulations of realistic scanner and number of protonts: ~1 mm accuracy
same on bone

## Slide 30: Real time anatomy and range measurement ?
MR-guided nitrogen-12-PET-verified proton therapy ?
Notes: MRI shows what is being treated, but does not reveal the proton beam.
MR-PET systems are on the market.
So why not combine MR-PET with proton therapy ?

## Slide 31: Real time anatomy and range measurement ?
ultrasound imaging and ionoacoustics ?
Kellnberger et al., Sci. Rep. 6(2016)29305
x
x
z
z
y
y
Notes: MRI shows what is being treated, but does not reveal the proton beam.
MR-PET systems are on the market.
So why not combine MR-PET with proton therapy ?

## Slide 32: Summary
How do protons behave inside the patient ?
  CT ➞ DECT
  DECT ➞ spectral CT ?
  proton radiography / CT
    (DE)CT calibration
    proton imaging
Was the dose delivered according to plan ?
  see what you treat: from X-rays to MR
  real time proton range measurement
    prompt gamma rays
    nitrogen-12 PET imaging
    ionoacoustics
Real time anatomy and range measurement
  MR + PET
  acoustics

## Slide 33: The people (now and recent past)
KVI-CART
Aleksandra Biegun, Sytze Brandenburg, Karol Brzezinski, Tom Buitenhuis, Peter Dendooven, Faruk Diblen, Oksana Kavatsyuk, Josbert Mulder, Francesco Olivari, Ikechi Ozoemelam, Joanne van Abbema, Emiel van der Graaf
UMCG Radiotherapy, Radiology
Marc-Jan van Goethem, Marcel Greuter, Arjen van der Schaaf, Erik Korevaar, Marleen Vonder, Rozemarijn Vliegenthart, Gert Jan Pelgrim