## Slide 1: QUAD development a TPC building block
Yevgen Bilevych, Klaus Desch, Harry van der Graaf, Markus Gruber, Fred Hartjes, Bas van der Heijden, Kevin Heijhof, Jochen Kaminski, Peter Kluit, Auke Korporaal, Cornelis Ligtenberg, Oscar van Petten, Gerhard Raven, Lucian Scharenberg, Tobias Schiffer, Sebastian Schmidt and Jan Timmermans
Nikhef
and
Physikalisches Institut Universität Bonn
9th Symposium on Large TPCs for Low-energy Rare Event Detection. Diderot University, December 12 - 14, 2018

## Slide 2: GridPix technology
Pixel chip with integrated Micromegas
=> InGrid
Grid set at negative voltage (300 – 600 V) to provide gas amplification
Very small pixel size (55 µm)
=> mostly detecting individual electrons
GridPix chip
55 µm

## Slide 3: GridPix: detecting individual electrons
Uses the complete information of the ionization profile
The best resolution that can be obtained with a gaseous detector
  dE/dx by single electron counting
Fine granularity enables rejecting background tracks and deltas in offline analysis
Measured track in test beam

## Slide 4: Pixel chip: TimePix3
256 x 256 pixels
55 x 55 µm pitch
=> 14.1 x 14.1 mm sensitive area
TDC with 600 MHz clock (1.7 ns)
Presently used in the data driven mode
  Each hit consists of the pixel address and time stamp of arrival time (TtA)
  Time over threshold (ToT) is added to register the signal amplitude
  => compensation for time walk
  Trigger (for t0) added to the data stream as an additional time stamp
  (Running in triggered mode is possible)
High power consumption
  ~ 1 A @ 2 V (2W), depending on hit rate
  => good cooling is important

## Slide 5: TimePix3 equipped with InGrid
Wafer post-processing at IZM
Aluminium grid (1 µm thick)
35 µm wide holes, 55 µm pitch
Supported by SU8 pillars 50 µm high
Grid surrounded by SU8 dyke (150 µm wide solid strip) for mechanical and HV stability
dyke
Cut view
50 µm

## Slide 6: Single chip test at Bonn (June 2017)
ELSA: 2.5 GeV electrons
Tracks referenced by Mimosa telescope
Gas: Ar/CF4/iC4H10 95/3/2 (T2K)
Ed = 280 V/cm
Detector with guard and field shaper
Published paper on 2017 testbeam: https://doi.org/10.1016/j.nima.2018.08.012
@Bonn
--------- Mimosa telescope ---------

## Slide 7: TimePix3 time walk correction
Time walk error: time of arrival depends on signal amplitude
Correction using Time over Threshold (ToT) as a measure for signal strength
Residual distribution improved
Higher order corrections did not yield further improvements
(Blum, Particle detection 2008)

## Slide 8: Single hit resolution in transverse direction
( expected)
B T

## Slide 9: Single hit resolution in transverse direction
( expected)
B T
Single hit resolution in pixel plane:
Depends on:
pixel size
Diffusion from fit
Note that:
A hit resolution of 250 µm is 25 µm for a 100-hit track ( 1 cm track length)
At 4 T ,

## Slide 10: Single hit resolution in longitudinal direction
( expected)
Single hit resolution in drift direction
Depends on
from fit
Diffusion from fit
The additional ToT cut (>0.60 µs) was applied to avoid large time walk errors
B T

## Slide 11: Deformations in pixel plane (Y) and drift direction (Z)
The RMS of the mean residuals is 7 µm in the pixel plane, and 21 µm (0.3 ns) in the drift direction Overall grid quality is very good
Simulations show an improvement in momentum resolution of a pixel TPC readout over a pad readout*
  Energy loss resolution (dE/dx) by truncated sum is 4.1% per meter*
* Report about Gridpix for the TPC of the ILD experiment at the ILC can be found in: C. Ligtenberg et al., "Performance of a GridPix TPC readout based on the Timepix3 chip”, Proceedings of the International Workshop on Future Linear Colliders (LCWS2018), Arlington, Texas, 22-26 October 2018. C18-10-22
Y
Z

## Slide 12: Covering large detection areas: QUAD
Guard
TPX3
TPX3
TPX3
TPX3
COld
CArrier
LV PS
Four-TimePix3 detector
All services (signal IO, LV power) are located under the detection surface
The area for connections was squeezed to the minimum
Detection surface can be extended by adding other QUADs at all 4 sides
  => no limit on detection area
39.6 x 28.38 mm

## Slide 13: Assembled QUAD
QUAD has an sensitive area of 68.9%
COCA dimensions 39.6 x 28.38 mm
COld CArrier
(COCA)
Guard electrode omitted
QUADs being produced in series
(14 QUADs)
Commercialization under study

## Slide 14: Single QUAD test in Bonn (October 2018)
ELSA: 2.5 GeV electrons
Tracks referenced by Mimosa telescope
QUAD sandwiched between Mimosa halves
  Largely improved track definition
Gas: Ar/CF4/iC4H10 95/3/2 (T2K)
Field cage
QUAD
@Bonn
Only preliminary results Analysis in progress

## Slide 15: QUAD time walk results
Time walk correction works well
Preliminary

## Slide 16: QUAD diffusion measurements
Preliminary

## Slide 17: QUAD edge deformations
Small deformations due to
  Dead zone between chips
  Grounded region between chips
May be corrected by proper guard electrode
Preliminary
Grounded region

## Slide 18: QUAD residuals in precision plane
guard
Outside telescope acceptance
Outside telescope acceptance
After applying edge corrections
The cut offs at left and right side are due to the telescope acceptance and beam coverage
Preliminary
Residuals well below 100 µm

## Slide 19: QUAD as a building block
Building blocks of 39.6 x 28.38 mm
Unlimited surface may be covered
Plug in QUADs into a cooled mounting plate
Push them from two sides to a mechanical reference
Sensitive area 68.9%
Cooling channels

## Slide 20: Mounting plate for 8 QUADs being assembled
Guard electrodes not yet installed

## Slide 21: Conclusions
Since 2017 three TimePix3 wafers were successfully equipped with an InGrid
A single chip GridPix detector from this production was reliably operated in a test beam in 2017
  The resolution is primarily limited by diffusion
  Systematic uncertainties are low: < 10 µm/cm in the pixel plane
  Data from a recent 2018 single QUAD test beam experiment has been presented
  Data quality and resolutions are similar to the single chip test beam results
Small edge deformations at the boundary between two chips are observed
  We can correct them but also study solutions with guard electrodes to reduce them
A production of 14 QUAD modules is almost finished
  QUADs are being installed in an 8-QUAD mounting plate

## Slide 22: BACKUP

## Slide 23: Maximizing active area
Using TimePix3 chip
Wirebond board squeezed to smallest width possible => 3 mm
Inactive zone (wirebond pads + electronics) => 2 mm
=> we have to reserve a 5 mm wide zone for connection + chip electronics
3 mm
2 mm
Sensitive area

## Slide 24: 2018 Testbeam track and hit selection

## Slide 25
New QUAD all hits

## Slide 26: QUAD’s tooling
Pickup tool for chip
with InGrid
Chip alignment jig
Assembly tool

## Slide 27: Chip aligment
Semi automatic under LabVIEW
Referring the bonding pads of the chips to the precise edges of the COCA
Aimed accuracy 20 µm in X,Y and Z

## Slide 28: QUAD’s active area
Each chip has 256 x 250 active pixels of 55 x 55 µm
  4 columns of pixels used for dykes
  => 4 x 193.6 = 774.4 mm2 in a QUAD
QUAD’s outer dimensions
  39.60 x 28.38 = 1123.8 mm2
=> QUAD has an active area of 68.9%
By using a possible future through via option, the active area can be increased to 86.4%

## Slide 29: Single chip detector
Very small active area coverage
  1.2% (whole structure)
  2.4% (PCB)
14 mm
14 mm