## Slide 1: Compass studies Aart Heijboer
Pressure sensor
LSP25HB
Accelerometer & magnetometer
LSM303AGR
Issues: In DOMs and on CLBs
High residuals (up to 10 deg errors when measuring rotations of 90,180,270 degrees)
Results seem badly reproducible

## Slide 2: Compass studies Aart Heijboer
Issues in DOMs and on CLBs (Catania & Athens)
High residuals (up to 10 deg errors when measuring rotations of 90,180,270 degrees)
Results badly reproducible.
Much worse than previous version which had LSM303D device.
LSM303AGR very popular device (e.g. there are several Arduino libraries, adafruit breakout board)
Hard to believe it is intrinsically bad / cannot work in environment with other electronics.
=> Investigate bare boards detail.

## Slide 3: Executive summary
turn on bit 1 of register 0x61 (offset correction)
and bit 7 of 0x60 (temperature compensation)
Implemented in CLB firmware & first tests on CLB+compass (Athens) seen to show good behaviour now.

## Slide 4
Turntable with 3 compass boards
2 lsm303AGR (‘v3’)
1 lsm303D (‘v2’) (not read out)
Raspberry pi 4 for
Motor control, DAQ, analysis (ROOT)
Motor with gearbox &
Optical rotation sensor / angle
measurement : 1 tick = 0.42o of turntable rotation

## Slide 5: offsets
Device output
Scale
M_So = 1.5+/-0.1 mGaus/LSB
“bridge” offset
can be measured by device
Temperature dependent.
True B-field
‘hard iron’ offset |M_TyOff|<60 mGaus – must be measured externally

## Slide 6: offsets
LSM303AGR has a mode where consecutive measurements return:
B + bridge-offset (after ‘set’ pulse)
B – bridge-offset (after ‘reset’ pulse)
Average = offset-corrected B-field
(still containing the ‘hard iron’ offset)
Difference = 2 x bridge-offset
Can correct for bridge offsets, but also measure and monitor their value
Device works with magnetized material who’s resistance
depends on B-field in some direction. The resistance is measured
by a Wheatstone Bridge.
Magnetization can be flipped.

## Slide 7
Heat gun
a.k.a. my wife’s hair dryer.

## Slide 8: Bridge offsets: temperature
Bridge offsets very dependent on temperature
50 units (20 deg) for 30 oC temp jup
Translates to error of 0.6 deg / oC
nb: these offsets are currently simply included in the DOM measurements.
Sensor 0

## Slide 9: Bridge offsets: temperature
Sensor 1
Bridge offsets very dependent on temperature
100 units (40 deg) for 30 oC jump.
Translates to error of 1.2 deg / oC
Offsets in z-direction even more sensitive
All we see in testing could be easily explained by T-dependence alone
  difference bare CLB -> inside DOM
  Bad reproducibility
Actually plotting absolute value
Hence bounce here

## Slide 10: Bridge offsets : drift
Clear temperature correlation, also visible
Here,
But also drift :
Remember 0.4o/LSB
=> Up to 6 degrees in 2 weeks.
Can continue to take data

## Slide 11: B-field (bridge-offset corrected): temperature
Without temperature compensation
With temperature compensation
Still some remaining effect
From datasheet
T-dependence of sensitivity M_So (in principle drops out of angle measurement, but still…. )

## Slide 12: Calibration & Angle measurement
Offset-correction and
Temp-compensation on
Make turns, fit ellipse
(x,y offset, x,y scale )
Offset < 60 mGaus (in spec)
Compare with turn-table
angle measurement (from rotary sensor)
Residuals ~1 deg
very correlated between the two devices
-> table bias or soft-iron contribution.
x,y offset
y-scale
x-scale
table angle
compass angle
table angle
Compass – table angle
table angle
compass angle
table angle
Compass – table angle

## Slide 13: Stability of calibration
Device 0
Device 1
Calibration constants stable within ~4 LSB = about 1.5 deg.
Some variations due to small change in sensor orientation
Some variations still unexplained. E.g. jump in y-center - accelerometer shows no change - both sensors see it

## Slide 14: Conclusions
  LSM303AGR board tested in stand-alone setup
  Device functioning according to specs
  Sensor offsets more complicated than assumed
    Temperature (and time)-dependent ‘bridge offset’ contribution must be corrected by device (via set/reset pulse)
    Remaining hard-iron offsets and scales to be measured in calibration as usual.
  Calibration offsets (hard-iron) after ‘bridge offset’ correction, stable to ~1.5 degrees.
  (probably partly due to setup)

## Slide 15: History
AHRS
  Complex data-processing on-board
  As much as possible disabled to be able to understand magnetometer data
LSM303D (v2)
  Bridge-offset correction always applied ‘under the hood’ by device
  datasheet : ‘set/reset pulse is automatically applied at each readout cycle’
  ‘hard iron’ offset correction applied by us ‘offline’
LSM303AGR (v3)
  Did not realize existence of time/temperature dependent ‘bridge offsets’
  Assumed ‘offset correction’ was not needed as we do it in calibration
  Datasheet not very clear, although it offers hints.

## Slide 16: Compass studies
LSM303AGR
  Suspect it is an AMR device : Anisotropic Magneto-Resistance
Two types of offset
  ‘bridge offset’:
    Internally measurable by device by flipping polarity of the magnetised material in the sensor.
    Temperature-dependent and possibly drifting
  remaining offset to be calibrated, aka ‘hard iron correction’ as usual
Bridge offset known to be temperature dependent

## Slide 17: Stability of calibration
First temp.
cycle
Variations
-origin unknown
-correlated!
Device 0
Device 1
Maybe setup is tilted?
Check accell.