## Slide 1: Pixel TPC gating and double grid studies
    Studies Pixel TPC Gating (a la ILD)
    Studies Pixel TPC with double grid; a grid on top of the GridPix
Yevgen Bilevych, Klaus Desch, Sander van Doesburg, Jean-Paul Fransen, Harry van der Graaf, Markus Gruber, Fred Hartjes, Bas van der Heijden, Kevin Heijhoff, Charles Ietswaard, Dimitri John, Jochen Kaminski, Peter Kluit, Naomi van der Kolk, Auke Korporaal, Cornelis Ligtenberg, Oscar van Petten, Gerhard Raven, Joop Rövekamp, Tobias Schiffer and Jan Timmermans

## Slide 2: Pixel TPC gating
LCWS19 presentation ILD gating GEM
Yumi Aoki (KEK)
Plots do not have the right geometry and deform …

## Slide 3
Transparancy Gate

## Slide 5: Modeling the Gating device
Making E field calculations for the device with approximate dimensions
  Make rectangular grid with 300 μm spacing and rim 30 μm. So optical transparancy 81%.
  The two planes are put 50 μm apart (a bit further than the Japanese Gate).
  There is a central voltage Vc applied that should correspond to the potential of the field without the gating device (here 250 V/cm). And a voltage difference dV that is controlled to switch the gate
  The upper part of the gate is at 10 mm and the lower at 9.95 mm from the bottom plate (ground). The anode is 550 mm away.
  Exact expressions for the E field of rectangular (plates) are used; by using mirror charges the ground is put at z = 0. By adding and subtracting charges a grid with holes is created.
Questions: how can we minimize distortions? How precise should we control voltages and dimensions (how flat should the gate be mounted)?

## Slide 6: Trajectories in Gating device
Gate closed -30 V
Gate open 1+1.5 V

## Slide 7: Trajectories in Gating device
Gate open 1+1.5 V
Gate open 0.1+1.5 V
To open gate stay close to nominal field (and potential)
Deformations are smaller < 20 μm for dV = 0.1 V
Note that large maximum deviations will lead to ExB distortions
Max deviation
shift trajectory
Max deviation
shift trajectory
Gate open 0.1+1.5 V

## Slide 8: Deformed Trajectories in Gating device
Gate open 0.1+1.5 V
mismatch 10 V middle gate
This means that the voltage of the Gate should be correct up to 5-10 V or the gate should be mounted with a precision of < 0.2-0.4 mm (250 V/cm). In that case the deviations remain less than < 20 μm
Max deviation
shift trajectory

## Slide 9: Conclusions: Gating device
In order to keep deviations/deformations smaller than 10-20 μm:
The potential difference dV over the gate should be kept very close to the nominal field (times distance) plus/minus 0.1 V
  This means that one should not tune the gate Voltage for the highest efficiency/transparancy, but use the above specification
The central voltage of the Gate Vc should be correct up to better than 5-10 V
The gate should be mounted with a precision of better than 0.2-0.4 mm in z
These specification can be achieved with the proposed gating device
  controlling of the potential difference to 0.1 V is feasable
  the tuning of the central voltage to better than 10 V is not difficult; the tuning of the Guard wires of the 8-Quad module was already a few Volts
It is usefull to mount a gate on the current module and measure deformations as a function of dV and Vc

## Slide 10: Pixel TPC with double grid to reduce the ion back flow
Question: can one reduce the Ion Back Flow of a GridPix detector?
We could design a GridPix detector using a double grid
The idea is that by creating two field regions, one with a medium field and one with a high field (our standard Grid Pix) one could reduce the ion backflow in two stages.
  The high field avalanche region has a measured IBF of 1.3%
  The aim is to reduce the IBF by another factor 100
  The second Grid replaces the Gating device and is always operational

## Slide 11
Design of a double Grid
High field
Intermediate Field
GridPix
Drift region
Second Grid
50 μm
e.g.
250 μm

## Slide 12
In (down) flow trajectories second Grid
Field ratio 10
Field ratio 240
Field ratio 40
Geometry: second grid at 0.250 mm (z); Cathode at 550 mm
Standard GridPix pitch 55 μm and hole 30 μm
Field ratio = mean Field (0-0.250 mm)/ mean Field z (2-550 mm)
Electron tracking without diffusion:
σ (rms) size of funnel (focussing E field) = 2.6-1.5-1.1 (Fr 10,40,240)

## Slide 13
Backflow (up) trajectories second Grid
Field ratio 10
Field ratio 240
Field ratio 40
Here the trajectories of ions from the bottom upwards are shown
The differences between the different field ratios are small

## Slide 14: Modeling of ion back flow
Modeling of the ion backflow is based on the measurements for a standard GridPix with FR 240. The ion backflow was measured to be 1.3%.
The Ion BackFlow is sensitive to the diffusion in the high field region.
The electron funnel size is 1.1 μm just from the field focussing.
For the GridPix (gap 50 μm) one expects an electron diffusion of 150-200 μm/√cm. This gives a smearing of the funnel of about 10-14 μm.
The calculated IBF corresponds to 1.3% for a smearing of 15 μm.
So this agrees reasonably well with expectations.
The performance IBF of the top grid.
What happens is that the back flowing ions that make it through the lower
grid that runs at FR240 will be flat distributed if one is 60-100 μm or more
above it.

## Slide 15: IBF of the second grid
The performance IBF of the second grid.
What happens is that the back flowing ions that make it through the lower grid that runs at FR240 will be flat distributed if one is 60-100 μm or more above it.
The field ratio should not be too high to avoid gas amplification in
the top grid. Therefore we leave out the FR240 point.
IBF (% ) FR 40 | FR 10
1.2 | 3.5

## Slide 16: Transparancy of the grid
Another important aspect is the electron transparancy.
Using the simulation this can be calculated.
It is important to choose a FR with a high (electron) transparancy
so with FR of 40 or higher.
Transparancy (%) FR 240 | FR 40 | FR 10
100. | 100. | 99.0

## Slide 17: Further reflections on the gap size
Another important aspect is diffusion that takes place in the
intermediate field region. For the T2K gas this can be at most 400 μm/√cm. The gap is the distance between the two grids.
So in case of a 1 mm gap there is a sizeable probability that the
neigbour pixels detect the avalanche.
So a smaller gap is preferable.
Smearing σ (μm) gap 1 mm | gap 250 μm | gap 60 μm
126 | 63 | 31

## Slide 18: Ion backflow for a double grid
Here calculations for the IBF of the two grids in case one has a total FR of
about 240 – normal GridPix operation. For the simulations the FR of the top
grid was put at 16. The lower Grid(Pix) was at FR 16 too. Total FR 256.
In order to reach IBF*Gain (2 103) below one has to choose a slightly
smaller hole size of 25 or 20 microns.
Ion backflow | Hole 30 μm | Hole 25 μm | Hole 20 μm
Top grid | 2.2% | 1.2% | 0.7%
GridPix | 5.5% | 2.8% | 1.7%
Total | 12 10-4 | 3 10-4 | 1 10-4
transparancy | 100% | 99.4% | 91.7%

## Slide 19: Conclusions: double grid
The Ion Back Flow can be significantly reduced by putting a grid with a
identical pitch and hole size on top of the Gridpix.
A device placed e.g. at 60-250 μm above the GridPix and ran with a
Field ratio of 16 (top) and 16 (lower) would do an excellent job.
The electron transparancy would be over 99 (91)% and the IBF would go
down from 1.3% to 3 (1) 10-4 for a hole size of 25 (20) μm.
This would solve the issue of IBF at CEPC and ILC.
We could do a test at Nikhef mounting this grid on top of the
Gridpix (holes 30 μm) and measure the electron transparancy and the
IBF and test the prediction on the previous slide.
It would be interesting to think about a post-processing step to
integrate the two grids.