## Slide 1: DESY ToT puzzles
Hardware related puzzles
What is the cause of the (largely) different ToT values for the chips?
  Is this due to the electronics of the TPX3 chip?
  Is it due to the gain differences? And due to the height varations of the pillars of the Grid
Here the plot for the ToT and mean ToT per chip for run 6969. After removing the large events where the quad lights up.

## Slide 2: DESY ToT puzzles
Is it the difference in the mean ToT due to the different thresholds in the chips?
No as shown in the plots below

## Slide 3: DESY ToT puzzles
Is it the difference due to a different electronics response for the different chips?
The testpulse data – where we inject a signal and scan the fine threshold and count the number of hits - show no clear different behaviour for chips 29, 30 or 31.

## Slide 4: DESY ToT puzzles
Is it the mean ToT constant over the surface of the three chips 29-31?
There are large systematic
variations of +- 2

## Slide 5: DESY ToT electronics calibration
ToT calibration plots from Kees for the 4 chips of the Bonn testbeam quad
The per pixel variations of the (electronics) ToT is 6.5% RMS
https://indico.cern.ch/event/267425/attachments/477859/661149/Timepix3_final.pdf
Chip 0
Chip 1

## Slide 6: DESY ToT electronics calibration
ToT calibration plots from Kees for the 4 chips of the Bonn testbeam quad
Systematics around 1000e is about 10%
Chip 2
Chip 3

## Slide 7: ToT conclusions
The per chip variations in the ToT calibration are not due to electronics gain differences (that can only produce up to 10% variations)
They are also not correlated to the position of the chip on the wafer (see talk by Jan 23 August https://indico.nikhef.nl/event/3207/)
They also not correlated to the thresholds
The low ToT values for chip 28-31 could be due to the power distribution of the multiplexer. The power is more critical for this quad as more channels go over threshold together (hurray events)
The variations between the chips and the variations over the chip of +- 2 counts could be due to variations of the charge up of the protection layer