## Slide 1: Measuring the Absolute efficiency of a Micro Channel Plate at Roma Tre and perspective of e- source
Alice Apponi, Alessandro Ruocco
Università Roma Tre and INFN Roma Tre
Gianluca Cavoto, Carlo Mariani
Università Sapienza and INFN Roma
Francesco Pandolfi, Ilaria Rago
INFN Roma
PTOLEMY meeting 06-07.10.2022

## Slide 2: MCP characterisation in the (30-900) eV energy range
Pulse Height Distribution (PHD) and Resolution (PHR)
Single- to multi-electron events discrimination
Absolute efficiency (e)

## Slide 3: Two-stage MCP
27 mm
14.5 mm
Hamamatsu MCP F1551-21S
Outer size = 27 mm
Effective diameter = 14.5 mm
Channel = 12 m
Open Area Ratio 60%
Two-stage chevron
Bias angle ~8°
Max MCP supply voltage 2.0 kV
MCP supply voltage 1800 V
∆Anode +100 V
Bias
angle

## Slide 4: We measure the absolute beam current or MCP output
e-gun
MCP
e-gun
Mono gun main feature:
Beam energy (30 - 900) eV
Beam current in fA range
Continuous beam
UH-Vacuum
~10-10 mbar

## Slide 5: Efficiency measurement: the method
TIME ~2h
Current scan in range [ ~5, ~280 ] fA
E2
Why?
MCP area = 14.5 mm
Avoid damages
Higher efficiency expected
Why?
FC hole = 3 mm
Slight beam shift during scan
Full current integration
FOCUSED beam: current scan on FC
LARGE beam: current scan on MCP
FOCUSED beam: current scan on FC
FOCUSED beam:
LARGE beam
Verify that the beam size variation modifies the current density only (not the total current!)
Check stability with current measurement before and after

## Slide 6: The beam size variation modifies the current density only
FOCUSED beam
LARGE beam
2D spatial scan on FC
Measurement time ~2h

## Slide 7: We have a good stability on current
Stability before-after difference
Accuracy 1% + 3 fA
Total uncertainty

## Slide 8: The efficiency is ~49% with no energy dependence
Efficiency defined as
Linear fit from min current to ~50 fA (300 kCounts/s)
Efficiency ~49% with no evident energy dependence within 30-900 eV
Fit result:
f(x) = q + mx
q = 0.489 ± 0.003
m = (1.8e-05 ± 5e-06) eV-1

## Slide 9: Pulse shape analysed with digital oscilloscope
~17 ns
∆A/A
N
1000 entries
Rhode&Schwartz Digital Oscilloscope RTA4004
Typical pulse shape
Reflection after ~17 ns always present
Reflection area is a fraction ~const of the main pulse (~36%)

## Slide 10: PHD: a weak energy resolution
Pulse Height
Discriminator threshold vt = 120 mV
Energy dependence of the Pulse Height Distribution (PHD)
Pulse Height Resolution
PHR = FWHM / Mean amplitude
@ 260 eV min PHR ~0.79

## Slide 11: Aligned and not aligned points
A further analysis of the pulse shape:
Look at the peak area-amplitude correlation
The majority of points are aligned
~1.5% are off-diagonal

## Slide 12: Aligned and not aligned points
~3 ns
~6 ns
~3 ns

## Slide 13: Third information: the peak width
Add a third information:
The peak width
Aligned points have ~same width
Off-diagonal are wider
Except for separated peaks
= 10 ns

## Slide 14: This behaviour does not depend on energy

## Slide 15: Not on energy nor on output rate
From linear fit on area vs amplitude

## Slide 16: The nature of the off-diagonal events
Two independent events in fixed time window :
= total events, = MCP rate
= 6 in this case
There is some instrumental effect
Which is energy dependent
20000 events
= 10 ns
= c/s
Previous slide example
Noff
t
𝜏

## Slide 17: To conclude
MCP characterisation with custom electron gun in range 30 - 900 eV at Roma Tre:
Absolute efficiency of the detector ~50% with no evident dependence on energy
Weak energy resolution observed in the Pulse Height Distribution (PHD) analysis
Further studies on pulse shape allows single- to multi-electron events discrimination, quantitative results require deeper analysis

## Slide 18: The key ingredients for measuring detector efficiency
Monogun at Roma Tre:
Monogun Main features | 
Energy range | 30 – 900 eV
Energy resolution | 45 meV
Current - range | 10 fA – 100 nA
Current - resolution | 200 aA
Spot size diam. | < 0.5mm
max count rate for MCP in single electron count mode: 300 kHz
minimum stable current from e-gun 10 fA
capability to measure the low current
emitted from e-gun (picoamm. Keysight B2987A) < 1fA

## Slide 19: Present and Future Experiments and Electron Sources
Specs (Roma Tre)
Monogun (Roma Tre)
Kimball (LNGS)
Available
e- source
Experiment
Detector efficiency
Ptolemy (e.m.)
Graphene transmission
Milan e-source
SPECS + new P.S.
Future
e- source
single count mode limit

## Slide 20: Thank you for your attention!!
Alice Apponi, Alessandro Ruocco
Università Roma Tre and INFN Roma Tre
Gianluca Cavoto, Carlo Mariani
Università Sapienza and INFN Roma
Francesco Pandolfi, Ilaria Rago
INFN Roma
PTOLEMY meeting 06-07.10.2022

## Slide 21: The key ingredients for measuring graphene transmission
Monogun at Roma Tre:
Monogun Main features | 
Energy range | 30 – 900 eV
Energy resolution | 45 meV
Current - range | 10 fA – 100 nA
Current - resolution | 200 aA
Spot size diam. | < 0.5mm
transmission through graphene measured with I=200 pA
minimum stable current from e-gun 10 fA
capability to measure the low current
emitted from e-gun (picoamm. Keysight B2987A) < 1fA

## Slide 22: Two beam sizes
Variation of the beam size:
Change polarisation of acceleration stage of e-gun
Spatial scan on Faraday Cup (FC)
Beam size = FWHM of the scan derivative
FWHM ~0.5 mm
FWHM ~3 mm
Derivate

## Slide 23
New test condition:
electron beam current measured with a Faraday Cup
picoamperometro Keysight B2987A
triaxial cable
common disconnected from ground
case of the picoamperometer connected to the same
ground of the F.C.
picoamperometer powered on battery

## Slide 24
Medium resolution e-gun: Test (I)
Stability
test with Faraday cup:
- range of current: 180 fA – 100 nA
- Stability vs time
- Beam dimension

## Slide 25: PHD: a weak energy resolution
Pulse Height
Discriminator threshold vt = 120 mV