## Slide 1
Precision Timing with Silicon Detectors
Silicon detectors with precision timing capabilities are playing an increasingly vital role in particle physics instrumentation. Among their many applications, pileup rejection and time-of-flight measurements at the Large Hadron Collider (LHC) stand out as prominent examples.
This presentation explores various silicon-based timing sensors, including silicon pad and pixel sensors, Low-Gain Avalanche Detectors (LGADs), Single-Photon Avalanche Diodes (SPADs), and Silicon Photomultipliers (SiPMs).
We examine the primary factors that influence time resolution—such as Landau fluctuations, electronic noise, variations in signal shape due to the pad response function, and gain fluctuations. The fundamental limits to timing precision imposed by these effects, along with practical considerations and constraints in real-world applications, will be discussed.
W. Riegler, CERN
NIKHEF, 12.6. 2025

## Slide 2
Planar sensors, 3D sensor, MAPs, LGADs, SPADs
W. Riegler and G. Aglieri Rinella, Point Charge Potential and Weighting field of a Pixel or Pad in a Plane Condenser, NIMA 767 (2014) 267-27
W. Riegler and G. Aglieri Rinella, Time resolution of silicon pixel sensors, JINST 12 (2017) P11017
W. Riegler, P. Windischhofer, Time resolution and efficiency of SPADs and SiPMs for photons and charged particles, NIMA 1003 (2021) 165265
P. Windischhofer, W. Riegler, The statistics of electron–hole avalanches, NIMA 1003 (2021) 165327
P. Windischhofer, W. Riegler, Passive quenching, signal shapes, and space charge effects in SPADs and SiPMs, NIMA 1045 (2023) 167627
W. Riegler, Time resolution limits in silicon sensors from Landau fluctuations and electronics noise, arXiv:2504.02570
W. Blum, W. Riegler, L. Rolandi, Particle Detection with Drift Chambers,
Springer 2008, ISBN: 978-3-540-76683-4
 Signal Processing, Noise, optimum filters etc
.

## Slide 3
Time resolution of ‘standard’ planar and 3D silicon detectors, Landau fluctuations
Signal processing and centroid time of a signal, noise, optimum filters
Contributions from ‘weighting fields’
Time resolution of MAPS sensors
Time resolution Low Gain Avalanche Diodes (LGADs)
Single Photon Avalanche Diodes (SPADs), Silicon Photomultipliers (SiPMs)
Outline

## Slide 4
Time resolution of ‘standard’ silicon sensors

## Slide 5
Intrinsic time resolution of a ‘large’ silicon pad detector
n+
n-
p+
-V
d
holes
electrons
x
E
In silicon sensors the signal edge is instantaneous (i.e. sub ps level)
acceleration of electrons to 107cm/s in vacuum is 0.14ps
passage of the particle through a 50um sensor takes 0.16ps
In Wire Chambers the electrons first have to move to the wires before an avalanche at the wire leads to an appreciable signal
 intrinsic resolution limit.
In RPCs the avalanche starts instantly, but it still takes some time until the signal reaches the threshold
 intrinsic resolution limit.
The intrinsic time resolution of a silicon sensor is infinite (sub ps).
 The time resolution in a planar silicon sensor without gain is a question of signal/noise/electronics and specifically the Landau fluctuations within the electronics integration time.
Example: d = 300um, Vdep = 58V, V=1.2 Vdep = 68V
d0 = 330um, 𝜏e = 5.6ns, 𝜏h = 16.8ns

## Slide 6
Total drift time of electrons and holes in the silicon sensor
At high electric fields, the drift-velocity saturates around 107 cm/s = 0.1μm/ps for both, electrons and holes.
In a 50um sensor at very high voltage the electrons and holes take around 500ps to move through the sensor.
At high fields (thin sensors, large voltage), the electric field in the sensor is close to constant, e and h velocities become saturated and the average signal becomes a triangle.
Example: d = 300um, Vdep = 58V, V=1.2 Vdep = 68V
d0 = 330um, 𝜏e = 5.6ns, 𝜏h = 16.8ns

## Slide 7
Electronics processing of a detector signal
Frontend delta response:
Corresponding transfer function:
f(t)
h(t)
-20 log |H(ω)|/|H(0)|
v(t)
tp=1ns fbw ~250MHz

## Slide 8
tp=0.5ns
tp=1ns
tp=2ns
tp=5ns
tp=10ns
f(t)
Electronics processing of a detector signal
v(t)
h(t)
If the peaking time tp becomes larger than the signal length, the peak of the output signal approaches the total charge. What else … ?

## Slide 9
Electronics processing of a detector signal
In case the electronics peaking time tp is longer than the signal duration T, the electronics output signal has
the same shape as the delta response
a pulse-height equal to the total charge of the signal
a ‘time displacement’ of this delta response by the centroid time tcog of the signal.
An amplifier that is ‘slower’ than the signal measures the centroid time of the signal
Signal duration of T i.e. f(t) = 0 for t>T
Electronics peaking time tp >> T
We are interested in times t>T

## Slide 10
There are many different ways to correct for this slewing effect
Constant Fraction discrimination
Standard discrimination using time over threshold to correct for pulse-height
Standard discrimination + pulseheight to correct for pulse-height
Standard discrimination + total charge to correct for pulse-height
Multiple sampling and ‘fitting’ the know signal shape
….
Electronics ‘slower’ than the detector signal, time slewing
Delta response shifted by tcog and scaled by Q
Signal normalized to same amplitude  time
‘time slewing’
 What is the centroid time resolution of a silicon sensor ?

## Slide 11
Energy deposit in silicon
5 clusters/μm
large cluster variation with long tail (1/n2)
resulting energy loss

## Slide 12
A high energy particle passing the sensor will experience primary interactions, with an average distance of around λ=0.21μm. In each interaction there is a probability pclu(n) to produce a cluster of n electrons. The probability p(n, Δz) have n e-h pairs in Δz is therefore
The probability to have n e-h pairs in the sensor of thickness d, p(n, d) is then
The cluster size distribution is pclu(n) is typically calculated using the PAI model.
Landau’s approach assumes a 1/E2 distribution for the energy transfer in accordance with Rutherford scattering, which leads to an 1/n2 distribution for the e-h pairs
 Landau distribution
Energy deposit in silicon
PAI model,
‘plasmon peak’

## Slide 13
Energy deposit in silicon
Most probable number of e-h pairs:
Full width of half maximum
3160 e-h pairs
14200 e-h pairs

## Slide 14
Single e-h pair
z=0.1d
z=0.5d
z=0.9d
Induced signal:

## Slide 15
Induced signal:
Centroid time:
Variance for uniform distribution of z:
Single e-h pair
Assuming

## Slide 16
L. Anderlini et al., Intrinsic time resolution of 3D-trench silicon pixels for charged particle detection. JINST 15, P09029, 2020.
D. Brundu et al., Accurate modelling of 3D-trench silicon sensor with enhanced timing performance and comparison with test beam measurements. JINST 16, P09028, 2021.
Total charge from the 200um sensor but timing characteristics from a 25um sensor !
10-20ps achievable and indeed achieved !
3D trench sensor

## Slide 17
Relating a single e-h pair to the Landau fluctuations
W. Riegler, Time resolution limits in silicon sensors from Landau fluctuations and electronics noise, arXiv:2504.02570

## Slide 18
Relating a single e-h pair to the Landau fluctuations
Landau theory

## Slide 19
Relating a single e-h pair to the Landau fluctuations
Assuming

## Slide 20
Center of gravity time resolution for silicon sensors
50um sensor, 10ps

## Slide 21
Leading edge discrimination
Leading edge discrimination on the signal that is normalized to the total deposited charge.
50um sensor, tp=0.5ns, threshold at 30% of the charge
5ps time resolution BUT this is unrealistic since S/N becomes too small  Noise will dominate !
50um sensor at 200V for different peaking times
and constant fraction discrimination T=0.8ns
W. Riegler, Time resolution limits in silicon sensors from Landau fluctuations and electronics noise, arXiv:2504.02570

## Slide 22
Noise
For a sensor that is represented by a capacitance, the noise is determined by the amplifier only. The amplifier noise can be characterized by the parallel and series noise power spectrum.
In case the parallel and series noise power spectra are ‘white’ we can formulate this as noise resistance Rs and Rp.
The level of this noise is a specification of the amplifier, there is no intrinsic noise in the sensor.
This noise level is very specific to the technology, but in general, lower noise requires more power (current through the input transistor etc.).
This places practical limits on the achievable noise level in systems with high granularity i.e. many channels on a small surface.
Rp
en2
in2
CD
Rs

## Slide 23
Noise and Optimum filters
Let us assume we have a signal f(t) with superimposed noise of a given noise power spectrum and we want to find the amplifier transfer function that maximizes
the signal to noise ratio for the best amplitude measurement, or
the slope to noise ratio for the best time measurement
 Theory of optimum filters

## Slide 24
A signal f(t) superimposed to a noise with a noise power spectrum w(ω) is processed by a filter with transfer function H(iω). The output signal and output noise r.m.s. are
The signal to noise ration is then
The expression has an upper limit given by the Schwartz inequality, which relates two complex valued functions according to
We insert
and have
The equal sign applies if ψ = c1 ϕ and we have the optimum filter transfer function
The output signal is
Optimum filter for best signal to noise ratio
f(t)  F(ω)
w(ω)

## Slide 25
We want to minimise the time resolution i.e. maximise the slope to noise
Following the steps from before we find an upper bound for the slope to noise ratio of
And the optimum transfer function as well as the output signal are
Since multiplication with iω refers to the derivative in the time domain we have the following relation:
The filter maximising the slope to noise ratio, i.e. the filter giving the best time resolution, is equal to the time derivative of the filter that maximises the signal to noise ratio !
Optimum filter for best slope to noise ratio

## Slide 26
Assuming a silicon sensor with negligible depletion voltage and saturated drift-velocity, the signal shape is a triangle:
For the noise we assume series noise and parallel noise together with a detector capacitance CD
The maximum slope to noise ratio is
If we neglect parallel noise we have
This filter is a-causal and can only be approximated in practise.
Optimum filter for best timing
Input signal
Delta response of optimum filter
Output signal
Rp
en2
in2
CD
Rs

## Slide 27
Preamp delta response
For the noise we assume series noise and parallel noise together with a detector capacitance CD
The output signal is
And the maximum slope evaluates to
To find the optimum peaking time and best time resolution we therefore have to minimise the function
Rp
en2
in2
CD
Realistic preamp transfer function
n=1
n=2
n=4

## Slide 28
Rp
en2
in2
CD
Realistic preamp transfer function
n=1
n=2
n=4
Neglecting parallel noise (which is a good approximation in most practical applications) the optimum electronics peaking time tp is between T and 1.5T.
The achieved time resolution is only about 30% worse than the best achievable one with the optimum filter !
This will give the smallest noise contribution to the time resolution.

## Slide 29
Combined effect of Landau fluctuations and noise
If the noise dominates, the optimum peaking time is about tp ~ 1-1.5T
If the noise is at a similar level to the contribution from Landau fluctuations, the optimum peaking time is still around the same level tp ~ 1-1.5T and the time resolution is equal to the the c.o.g. time resolution of the silicon sensor.
If the noise is subdominant, shorter peaking times can be improve the time resolution, but with the divergence of the noise contribution at low tp this will not reach far below the c.o.g. time resolution.
In short: tp ~ T and σ = σc.o.g. ⊕ σnoise give a good order of magnitude for the achievable time resolution
Landau Fluctuations
Landau Fluctuations ⊕ Noise
Noise
scales with
enCD
c.o.g. time resolution

## Slide 30
In case noise is not dominant, the time resolution of a silicon sensor is somewhere between zero and the c.o.g. time resolution.
For a silicon detector signal with total duration T, the series noise contribution is minimised for amplifiers of tp ~ T.
For very low noise a time resolution better than the c.o.g. time resolution can be achieved, but the divergence of the noise contribution for tp  0 will limit this improvement.
Time resolution of ‘standard’ silicon sensors

## Slide 31
Effects of finite Pixel Size
Up to now we discussed the situation of a very large pad of size w>>d i.e. and ‘infinitely extended parallel plate geometry’.

## Slide 32
A uniform charge distribution
Charge moving towards the strip:
Charge moving away from the strip:
Total induced current on the strip
For very wide strips (parallel plate geometry) the induced charge from the movement of +λ and –λ is the same and equal to λd/2.
For small strips, the fraction of the signal due to the charges moving towards the strip increases.
Weighting field of a strip in a parallel plate geometry
w>>d
w=d
w=d/2
+λ
-λ
v1 = 3v2 e.g. electrons moving to the strip and holes moving away from the strip
v1
v2
d

## Slide 33
Weighting field of a strip in a parallel plate geometry
w>>d
w=d
w=d/2
+λ
-λ
v1 = v2/3 e.g. holes moving to the strip and electrons moving away from the strip
v1
v2
d
Different positions of the particle inside the pixel will lead to different pulse-shapes.
The varying signal shape will lead to variations in the c.o.g. time and therefore affect the time resolution.
This is also called the ‘weighting field effect’.

## Slide 34
Weighting field fluctuations
w/d = pixel size/detector thickness
This weighting field effect is strongly correlated with the Landau fluctuations.
The structure of the expression for the resolution is however similar to the one before.

## Slide 35
Weighting field fluctuations
Because electrons and holes have different velocities, it makes a significant difference whether the electrons or the holes move to the pixel.
For higher fields (thinner sensors) this difference will decrease.
The dependence on the different parameters is complex.
These fluctuations can dominate the time resolution !

## Slide 36
Time resolution of ‘standard’ silicon sensors
Good time resolution demands thin sensors.
Thin sensors give small charge and large capacitance i.e. un-favorable S/N and k/N.
Capacitance can be reduced by making the pixels small.
If the pixel size is in the same order as the sensor thickness, the weighting field fluctuations start to dominate … and there will be many channels …
… between a rock and a hard place …
Sensors with internal gain to overcome the noise limit (like gas detectors !)
Turn the by sensor 90 degrees and realise a parallel plate geometry in 3D !

## Slide 38
Time resolution of MAPS with small readout diodes
One flavour of MAPS sensors uses a small ‘collection’ electrode to minimise the capacitance and therefore the noise  E.g. ALPIDE
Since the ‘weighting field’ of the ‘readout electrode’ is concentrated around the diode, the signal consists to a good approximation of a sequence of delta pulses for each electron arriving at the diode.
Arrival time distribution.
The time resolution will be somewhere between the arrival time statistics of the first electron and the center of gravity time of all electron arrival times.
The values depend entirely on the geometry and electric field distribution in the sensor and MC simulations are used for evaluation.
GARFIELD++ simulation
Jan Hasenbichler

## Slide 39
FASTPIX: a monolithic CMOS Sensor with <200ps timing
WORKSHOP ON PICO-SECOND TIMING DETECTORS FOR PHYSICS
9–11 Sep 2021
University of Zurich
Eric Buschmann
https://indico.cern.ch/event/861104/contributions/4503032/
< 200ps achieved with a MAPS sensor !

## Slide 41
Time resolution of LGADs

## Slide 42
Avalanche Photo Diode, Low Gain Avalanche Diode
Idea goes back to the 1960ies.
A high field region is implemented in a silicon sensor by doping.
Electrons will produce an avalanche in this high field region.
The high field region is implemented by doping and related ’space-charge’ in the volume.
The sensor is operated in a region where there is electron multiplication but not yet hole multiplication.
This allows to have thin sensors (high field, short signal) but still have enough signal charge to overcome the limitation from noise.
For higher fields  electron+hole multiplication  avalanche divergence  quench resistor  SiPM
Electron (alpha) and hole (beta)
multiplication coefficient in silicon

## Slide 43
AvalanchePhotoDiode
holes
electrons
Gain holes
Total drift time for a given sensor thickness
A high field region is implemented in a silicon sensor by doping.
Electrons will produce an avalanche in this high field region.

## Slide 44
LGAD
z=0.1d
z=0.5d
z=0.9d

## Slide 45
Induced signal:
Centroid time:
Variance for uniform distribution of z:
Variance for Landau fluctuations for G>10:
LGAD

## Slide 46
LGAD
LGAD centroid time resolution
Standard silicon sensor centroid time resolution
50um sensor at 200V: 30ps
200um sensor at 200V: 140ps
50um sensor at 200V: 13ps
200um sensor at 200V: 70ps
The c.o.g. time resolution of LGADs is worse than the one for standard silicon sensors due to the very different signal characteristics – essentially an electron arrival time distribution.
Of course – the fact that the signal is larger by a factor 10-15 allows much more relaxed noise requirements, larger pixels etc. …

## Slide 47
LGAD Leading edge discrimination
Standard silicon sensor
For LGADs the time resolution does not improve over the centroid time resolution when using leading edge discrimination on the signal normalized to the total deposited charge.
While for standard silicon sensors, the leading edge of the signal is proportional to (equal to) the total amount of charge in the sensor, for the LGADs there is no relation between the leading edge and the total charge deposit.

## Slide 48
Weighting field fluctuations
w/d = pixel size/detector thickness
The effect of finite pixel size and varying signal shape with the weighting field is de-correlated from the Landau fluctuations, in contrast to the standard silicon sensors.
The expression very basic and universal !
The Landau fluctuations are related to the total electron drift-time
The weighting field fluctuations are related to the total hole drift time.
These expressions are valid for perpendicular tracks
T1=d/ve = total electron drift time
T2=d/vh = total hole drift time

## Slide 49
Time resolution of LGADs
Centroid time resolution is worse when compared to standard silicon sensors.
The big advantage is of course the much increased signal charge and therefore much relaxed noise specification, which allows very thin sensors for excellent time resolution as well as large pads.
The resolution cannot be significantly improved when using faster amplifiers. The centroid time resolution is essentially THE time resolution of the LGAD sensor.
Weighting field fluctuations and Landau fluctuations are de-correlated.

## Slide 50
Time resolution of Single Photon Avalanche Diodes
(SPADs)

## Slide 51
Single Photon Avalanche Diodes
d
x=0
x=d
elec.
hole
For very large electric fields, also the holes start to contribute to the avalanche.
α dx = probability for an electron to produce an additional e-h pair along a distance dx
β dx = probability for a hole to produce an additional e-h pair along a distance dx

## Slide 52
SPADs Breakdown Condition
α dx = probability for an electron to produce an additional e-h pair along a distance dx
β dx = probability for a hole to produce an additional e-h pair along a distance dx
d
x=0
x=d
elec.
hole

## Slide 53
Solution for the average Signal
A number of ne0 electrons and nh0 holes are placed at x=x0 at time t=0.
What is the average distribution of electrons ne(x,t) and holes nh(x,t) in the sensor at time t ?
What is the total number of electron Ne(t) and holes Nh(t) in the sensor at time t ?
Average signal I(t) = EW [ Ne(t) ve + Nh(t) vh ]
λ is must satisfy the ‘Eigenvalue’ equation:
The equation has a finite number >=1 of real solutions and an infinite number of complex valued solutions.
 For long times, the largest real Eigenvalue λ dominates.

## Slide 54
Solution for the average Signal

## Slide 55
Solution for the average Signal

## Slide 56
Quenching
The average signal will increase until the number of charges that are deposited on the capacitor lowers the voltage across the diode below the breakdown voltage.
The ‘quench’ resistor will determine the rate of recharging of the SPAD.
The voltage drop represents the SPAD signal.
This is the same principle as the Geiger counter, where the divergence of the avalanche is due to electrons and photons (produced in the avalanche and producing more ionization electrons).
C
R
ΔV=Q/C
SPAD
V
Rutherford and Geiger 1908

## Slide 57
Quenching
P. Windischhofer, W. Riegler, Passive quenching, signal shapes, and space charge effects in SPADs and SiPMs, NIMA 1045 (2023) 167627

## Slide 58
SPAD time resolution
Time measurement by setting a threshold to the diverging avalanche signal.
First investigate a SPAD without boundaries:
Exact solution exists
Features are similar to the (much more complicated) SPAD with boundaries

## Slide 59
Fluctuations of the avalanche, SPAD without boundary
Time response function when placing a threshold of n electrons to the signal.
Approximation for ‘large thresholds’
Time resolution is independent of the threshold
Time resolution is inversely proportional to the ‘risetime’ of the signal
For a single e-h pair, A=1 and = π/√6 = 1.28=O(1)

## Slide 60
Avalanche fluctuations, SPAD with boundary
Equations for the average electron and hole densities
Equations for second moment of the electron and hole densities
… surprisingly complex.
Solution gives important insights:

## Slide 61
Time resolution of SPAD
γ determines the exponential growth of the avalanche for long times
1/γv* determines the time resolution when setting a ‘high’ threshold to the signal !
The coefficient is O(1)  1/γv* IS the approximate time resolition of a SPAD
The electron and hole densities become fully correlated over the timescale of 1/γv* i.e. there are no more fluctuations. The fluctuations take place in the very beginning of the avalanches – from there on there is only deterministic grow of the avalanche.

## Slide 62
Time resolution aSPAD
Approximate r.m.s. SPAD time resolution 1/γv*
<10ps is in the cards …

## Slide 63
Other contributions to the time resolution of SPADs
Up to now we assumed a given number of electrons and holes at a given position x0.
 Only avalanche fluctuations are contributing to the time resolution.
For the measurement of single photons, the conversion point in the gain layer can vary.
For the measurement of charged particles, the Landau fluctuations of the charge deposit will contribute to the time resolution.
Conclusion (see paper): for all these scenarios, the time resolution stays around with c0 in the order of 1-3.
Photon interacting at a random position inside the gain layer
MIP
c0
c0
Expect excellent time resolution for MIPs

## Slide 64
Other contributions to the time resolution of SPADs
All the previous considerations on the SPAD time resolution are correct if we assume that there is at least one e-h pair produced in the center of the high field region.
A SPAD array can of course not realise this high field region over an extended area, there are edge effects.
At the saturated drift-velocity of 107 cm/s it takes an electron 10ps to move 1um.
The variation of the conversion point outside the high field region will therefore represent the major limit of time resolution for a single photon or a MIP.
One recent attempt for MIP timing uses a Cherenkov radiator on front of the SiPM.
 1 MIP produced man photons and there is a high chance that at least one will interact in the high field region.

## Slide 65
https://cds.cern.ch/record/1500583
https://garfieldpp.web.cern.ch/garfieldpp/
https://garfield.web.cern.ch/garfield/
Garfield and Garfield++

## Slide 66
https://doi.org/10.1016/j.nima.2005.08.064
https://magboltz.web.cern.ch/magboltz/
Magboltz and Heed

## Slide 67
Timing simulations for silicon with GARFIELD++
LGAD
Interfacing to GEANT:
ALPIDE

## Slide 68
Thank You !

## Slide 73
3D sensor realising a parallel plate geometry, TimeSPOT
L. Anderlini et al., Intrinsic time resolution of 3D-trench silicon pixels for charged particle detection. JINST 15, P09029, 2020.
D. Brundu et al., Accurate modelling of 3D-trench silicon sensor with enhanced timing performance and comparison with test beam measurements. JINST 16, P09028, 2021.
For a perfectly perpendicular track:
‘box’ signals from electrons and holes.
Landau fluctuations affect just the total pulse-height, which can be corrected.
c.o.g. time of the signal
Variance of the c.o.g. Time for uniform distribution of tracks
d
Total charge from the 200um sensor but timing characteristics from a 25um sensor !
10-20ps achievable and indeed achieved !
Evaluation of the formula

## Slide 74
FASTPIX: a monolithic CMOS Sensor with <200ps timing
WORKSHOP ON PICO-SECOND TIMING DETECTORS FOR PHYSICS
9–11 Sep 2021
University of Zurich
Eric Buschmann
https://indico.cern.ch/event/861104/contributions/4503032/
< 200ps achieved with a MAPS sensor !

## Slide 75
Time resolution of ‘standard’ planar silicon detectors
Signal processing and centroid time of a signal
Contribution from noise, optimum filters
Contributions from ‘weighting fields’
Time resolution of MAPS sensors
Time resolution Low Gain Avalanche Diodes (LGADS)
Single Photon Avalanche Diodes (SPADs), Silicon Photomultipliers (SiPMs)
SPAD time resolution and efficiency for photons and charged particles
Outline

## Slide 76
Efficiency of SPADs
Applying a voltage below the breakdown voltage, the avalanche will never diverge and a finite number of electrons and holes will be produced in the avalanche.
Applying a voltage above the breakdown voltage, some avalanches will diverge and some avalanches will only produce a finite amount of charge, typically much too small to be detected.
A SPAD is considered efficient only in case an avalanche diverges.
Probability Pe(x) for single electron at position x to produces a diverging avalanche and
Probability Ph(x) for single electron at position x to produces a diverging avalanche

## Slide 77
Efficiency of SPADs
Solution for a SPAD of thickness d
d
x=0
x=d
elec.
hole
Efficiencies for single electrons and holes.

## Slide 78
Efficiency of SPADs for MIPS
Solution for a SPAD of thickness d
Once a MIP leaves a cluster in a SPAD that is biased above the breakdown voltage, there is a quite high probability that a diverging avalanche is produced.
Efficiency ~ 1-exp[-d/λ]

## Slide 79
An e-h pair is produced at position z.
The electron arrives at z=0 at time T=z/v1.
The electron multiplies in the high field in the layer at z=0 (infinitely thin).
The holes move back to z=d inducing the dominant part of the signal (all in this approximation).
Centroid time resolution for standard silicon sensor:
Centroid time resolution for LGAD
The signal is now defined by the arrival time distribution of the electrons at the gain layer.
LGAD